Table Specification of subject of performance

„Center for the development of carbon canisters specification EURO 7 – Climatic Chamber with humidity control”

The contracting authority specifies to the tenderer special technical conditions for the project. By their conditions the contracting authority defines the characteristics of the inquired performance which the supplier has to comply with in their project realization. The tenderer states in the technical conditions whether they can provide performance in compliance with the requirements described in the column „Meeting the requirements“, they select an option and they state „YES“ in case they can perform according to this requirement and „NO“ in case they cannot meet the requirement. If the suppliers states „NO“, be it in a single entry only, they will be excluded from the tender or not being able to meet all requirements. If the supplier states „YES“ and it will be found out that the offered goods do not comply with this requirement, they can be excluded for non-compliance and for breaking the tender terms. If the tenderer fails to choose „YES“ or „NO“ where required, they can be excluded from the tenderer for not meeting the tender terms. In column „Parameter – offered value“ they state a specific parameter (in units as required) or some further specification offered performance is considered to correspond exactly to the requirement of the contracting authority specific in column „Required parameter“. The tenderer shall fill in the technical conditions according to the instructions state in them including the kind and type of goods. On filling in these kinds and types off goods the suppliers is bound to deliver them and their list will form an Annex to the Purchase Contract. This means that the supplier will be bound to deliver exactly what they have stated in their quote.

The supplier will supply new and unused Climatic Chamber with humidity control.

| Product trade name: | (Tenderer fills in the product trade name) |  |  |
| --- | --- | --- | --- |
| Required parameter | Parameter – required value | Meeting the requirements | Parameter – offered value |
| External dimensions of the chamber (HxWxD) | max. 2100x1200x2000 mm | YES / NO |  |
| Test area volume | Min. 500 L | YES / NO |  |
| Test area dimensions (HxWxD) | min. 900x800x700 mm | YES / NO |  |
| Test area construction | Welded from stainless steel | YES / NO |  |
| Test temperature range | min. -40 °C to + 150 °C | YES / NO |  |
| Test humidity range | min. 10 to 98 % | YES / NO |  |
| Dew point temperature range | min -3°C to +94°C | YES / NO |  |
| Cooling temperature change rate (according to IEC 60068-3-5) | min 3 K/min | YES / NO |  |
| Rate of change in heating temperature (according to IEC standard 60068-3-5) | min. 4 K/min | YES / NO |  |
| Window in the door | Required | YES / NO |  |
| Illumination of the test area | Required | YES / NO |  |
| Passages through the test area | min. 2x Ø50 mm incl. sealing of passages | YES / NO |  |
| The refrigerant used must comply with EC Regulation No 517/2014 and all cooling circuits must have a refrigerant of less than 2500 GWP | Required | YES / NO |  |
| PC software for remote control of the chamber (starting and shutting down the chamber, possibility of creating temperature and humidity profiles, monitor of chamber running, archiving of results). | Required | YES / NO |  |
| Automatic demineralised water replenishment | Required | YES / NO |  |
| Interface | Ethernet, USB | YES / NO |  |
| Potential-free contacts to turn off the test sample | min. 4pc, 24V and 0.5A | YES / NO |  |
| Digital inputs and Digital outputs for test sample control | min. 4 inputs and 4 outputs | YES / NO |  |
| Temperature protection of the sample by independent temperature sensor (according to EN 60 519-2) | Required | YES / NO |  |
| Power consumption | max. 10 kW | YES / NO |  |
| Noise at 1m according to ISO 11201 | max. 65 dB | YES / NO |  |
| Weight | max. 1000 kg | YES / NO |  |
| Accredited temperature calibration according to EN ISO/IEC 17025 in 2 points temperature and 2 points humidity after commissioning on site incl. issuing a calibration protocol | Required | YES / NO |  |
| EC/EU declaration of conformity | Required | YES / NO |  |
| Test report generation | Automatic | YES / NO |  |
| Operation Manual | Required | YES / NO |  |
| System Schematic | Required | YES / NO |  |
| Operator training | Onsite training required | YES / NO |  |

I (we) below signed       hereby honestly declare that the data stated above are true and genuine and that the supplier       shall deliver the goods exactly as stated in the technical and business conditions in their quote in case they are chosen in the subject public tender by the contracting authority.

In     on

_________________________________

Name and fiction of the authorized person

Stamp and signature of the authorized person