Table Specification of subject of performance

„Center for the development of carbon canisters specification EURO 7 - Sealed chamber for measuring hydrocarbon emissions”

The contracting authority specifies to the tenderer special technical conditions for the project. By their conditions the contracting authority defines the characteristics of the inquired performance which the supplier has to comply with in their project realization. The tenderer states in the technical conditions whether they can provide performance in compliance with the requirements described in the column „Meeting the requirements“, they select an option and they state „YES“ in case they can perform according to this requirement and „NO“ in case they cannot meet the requirement. If the suppliers states „NO“, be it in a single entry only, they will be excluded from the tender or not being able to meet all requirements. If the supplier states „YES“ and it will be found out that the offered goods do not comply with this requirement, they can be excluded for non-compliance and for breaking the tender terms. If the tenderer fails to choose „YES“ or „NO“ where required, they can be excluded from the tenderer for not meeting the tender terms. In column „Parameter – offered value“ they state a specific parameter (in units as required) or some further specification offered performance is considered to correspond exactly to the requirement of the contracting authority specific in column „Required parameter“. The tenderer shall fill in the technical conditions according to the instructions state in them including the kind and type of goods. On filling in these kinds and types off goods the suppliers is bound to deliver them and their list will form an Annex to the Purchase Contract. This means that the supplier will be bound to deliver exactly what they have stated in their quote.

The supplier will supply new and unused Sealed chamber for measuring hydrocarbon emissions.

| Product trade name: | (Tenderer fills in the product trade name) |  |  |
| --- | --- | --- | --- |
| Required parameter | Parameter – required value | Meeting the requirements | Parameter – offered value |
| Equipment to run following test methods: | 1-day, 2-day, 3-day Diurnal Bleed Emissions of hydrocarbons with various options for test temperature profile templates. (for example. US CARB/EPA LEV3, EU6d, China6). | YES / NO |  |
|  | Measurement of hydrocarbon permeation for parts placed in microSHED chambers for various temperature profiles. | YES / NO |  |
| Thermal chamber temperature range | Min. range 15 - 50 °C | YES / NO |  |
| Thermal Chamber air mixing | per US CAR LEV3 BETP | YES / NO |  |
| Thermal Chamber temperature measurement | per US CAR LEV3 BETP | YES / NO |  |
| Equipment construction | Main chamber with 3 microSHEDs. Main chamber configured only as a thermal control chamber without emission measurement. MicroSHEDs placed on wheels for easy manipulation. | YES / NO |  |
| Chamber construction material | Walls from stainless steel. Other materials approved for emission testing | YES / NO |  |
| Thermal Chamber Size Estimated internal size needed (2560mm X 4570mm X 2550mm) | Chamber must be designed to hold: 3 x microSHED with internal dimensions 900mm X 900mm X 920mm (WxLxH) placed along one side. 3 x Fuel tanks with external dimensions 860mm X 1380mm X 550mm (WxLxH). Fuel tanks on trolley with two floors. Two tanks on bottom floor and one tank on top. Trolley placed along opposite side of the chamber. | YES / NO |  |
| Access to thermal chamber | Walk in with ramp for trolley | YES / NO |  |
| FID | 3 separate FIDs. One for each microSHED | YES / NO |  |
| FID calibration | Automatic | YES / NO |  |
| Retention/Recovery | Automatic program with built in propane injection | YES / NO |  |
| Chamber cleaning | Built in program with bake off temp. 50°C | YES / NO |  |
| Leak test | Automatic for each microSHED | YES / NO |  |
| EC/EU declaration of conformity | required | YES / NO |  |
| ATEX declaration | required | YES / NO |  |
| Data logging frequency | 1 Hz | YES / NO |  |
| Test report generation | Automatic | YES / NO |  |
| Remote access | Required | YES / NO |  |
| Operation Manual | Required | YES / NO |  |
| System Schematic | Required | YES / NO |  |
| Operator training | Onsite training required | YES / NO |  |
| Verification testing | Required on site | YES / NO |  |

I (we) below signed       hereby honestly declare that the data stated above are true and genuine and that the supplier       shall deliver the goods exactly as stated in the technical and business conditions in their quote in case they are chosen in the subject public tender by the contracting authority.

In     on

_________________________________

Name and fiction of the authorized person

Stamp and signature of the authorized person