Table Specification of subject of performance

„Center for the development of carbon canisters specification EURO 7 – Station for cyclic loading of canisters with Butane and Gas”

The contracting authority specifies to the tenderer special technical conditions for the project. By their conditions the contracting authority defines the characteristics of the inquired performance which the supplier has to comply with in their project realization. The tenderer states in the technical conditions whether they can provide performance in compliance with the requirements described in the column „Meeting the requirements“, they select an option and they state „YES“ in case they can perform according to this requirement and „NO“ in case they cannot meet the requirement. If the suppliers states „NO“, be it in a single entry only, they will be excluded from the tender or not being able to meet all requirements. If the supplier states „YES“ and it will be found out that the offered goods do not comply with this requirement, they can be excluded for non-compliance and for breaking the tender terms. If the tenderer fails to choose „YES“ or „NO“ where required, they can be excluded from the tenderer for not meeting the tender terms. In column „Parameter – offered value“ they state a specific parameter (in units as required) or some further specification offered performance is considered to correspond exactly to the requirement of the contracting authority specific in column „Required parameter“. The tenderer shall fill in the technical conditions according to the instructions state in them including the kind and type of goods. On filling in these kinds and types off goods the suppliers is bound to deliver them and their list will form an Annex to the Purchase Contract. This means that the supplier will be bound to deliver exactly what they have stated in their quote.

The supplier will supply new and unused Station for cyclic loading of canisters with Butane and Gas.

| Product trade name: | (Tenderer fills in the product trade name) |  |  |
| --- | --- | --- | --- |
| Required parameter | Parameter – required value | Meeting the requirements | Parameter – offered value |
| Number of Canister loading stations | 3 | YES / NO |  |
| Number of scales per loading station | 2 (one scale for tested and one for reference canister). | YES / NO |  |
| Scale bay size | Must fit canister of max. dimensions 500x500x500mm with enough room for connections. | YES / NO |  |
| Equipment construction | Canister loading stations located along one side of the wall with gas generating equipment on the other side of the wall. Canister loading stations connected thru the wall with gas generating equipment with connections as short as possible. | YES / NO |  |
| Butane flow rate range (each station) | Up to 100 g/hr | YES / NO |  |
| Nitrogen flow | Needs to maintain 50:50 ± 15% concentration with Butane or Gasoline | YES / NO |  |
| Normal gas flow rate range (each station) | Up to 200 g/hr | YES / NO |  |
| ORVR gas flow rate range (each station) | Up to 3600 g/hr | YES / NO |  |
| Purge air flow rate range (each station) | Up to 100 slpm. Flow to be generated by vacuum pump. | YES / NO |  |
| Vapor generation method | Nitrogen Gas and/or Heat | YES / NO |  |
| Automatic fuel change | Required | YES / NO |  |
| Connections | 200L barrel for new and used fuel. Transfer pumps and level sensors to be included. | YES / NO |  |
| The tests to be in compliance with the following regulations: |  |  |  |
| CARB | California Evaporative Emission Standards and Test Procedures for 2001 and subsequent Model Motor Vehicles, Issue 05.08.1999 – last amendment October 17, 2007 | YES / NO |  |
| EPA | Code of Federal Regulations 40 Part 86, Issue 2014 | YES / NO |  |
| EU | Annex to the Commission Regulation xxx/2016 amending Commission Regulation (EC) No 692/2008 as regards the methodology for the determination for evaporative emission (Type 4 test) | YES / NO |  |
| China 6 | GB18352.6—2016, Annex I | YES / NO |  |
| Equipment to run following procedures (tests): |  |  |  |
| Purge only |  | YES / NO |  |
| Working Capacity Determination (breakthrough) | CARB Part III D 3.3.7(a) EPA §86.132-96(j)(1) | YES / NO |  |
| Pre-Conditioning and stabilization | CARB Part III D 3.3.4 EPA §86.132-96(h)(1) | YES / NO |  |
| Canister Aging/ Canister WC Determination | EU6 | YES / NO |  |
| GWC |  | YES / NO |  |
| China BWC | China 6 | YES / NO |  |
| Leak Check |  | YES / NO |  |
| Option to configure additional test procedures | Required | YES / NO |  |
| Humidity control system for purge air supply | 50 ± 25 grains of water per pound of dry air or 7.13 grams of water per kg of air, equivalent to a relative moisture content of 50% at 21°C with a tolerance range from 25 to 75% at 21°C | YES / NO |  |
| Dynamic purge flow | Required | YES / NO |  |
| Purge air temperature | Up to 100F (38°C) | YES / NO |  |
| Purge air temp. measurement | Temp. to be measured close to Canister inlet to ensure correct temp during test cycle. | YES / NO |  |
| Independent LEL sensor for gas leak monitoring | Required | YES / NO |  |
| Computer and control software | Required | YES / NO |  |
| BWC Repeatability test done in delivery location. 13 cycle BWC at 15g/hr and 40g/hr using 1 liter canister with approx. 60g working capacity. | MFCs and Scales to correlate within 2 grams after stabilization. Station to station correlation to be within 2 grams. | YES / NO |  |
| GWC Repeatability test done in delivery location. 13 cycle GWC at 15g/hr and 40g/hr using 1 liter canister with approx. 60g working capacity. | MFCs and Scales to correlate within 2 grams after stabilization. Station to station correlation to be within 2 grams. | YES / NO |  |
| EC/EU declaration of conformity | Required | YES / NO |  |
| ATEX declaration | Required. Equipment will be operated in a non-hazard area. | YES / NO |  |
| Data logging frequency | 1 Hz | YES / NO |  |
| Test report generation | Automatic | YES / NO |  |
| Remote access | Required | YES / NO |  |
| Operation Manual | Required | YES / NO |  |
| System Schematic | Required | YES / NO |  |
| Operator training | Onsite training required | YES / NO |  |

I (we) below signed       hereby honestly declare that the data stated above are true and genuine and that the supplier       shall deliver the goods exactly as stated in the technical and business conditions in their quote in case they are chosen in the subject public tender by the contracting authority.

In     on

_________________________________

Name and fiction of the authorized person

Stamp and signature of the authorized person