Table Specification of subject of performance

„Center for the development of carbon canisters specification EURO 7 – Vibration table”

The contracting authority specifies to the tenderer special technical conditions for the project. By their conditions the contracting authority defines the characteristics of the inquired performance which the supplier has to comply with in their project realization. The tenderer states in the technical conditions whether they can provide performance in compliance with the requirements described in the column „Meeting the requirements“, they select an option and they state „YES“ in case they can perform according to this requirement and „NO“ in case they cannot meet the requirement. If the suppliers states „NO“, be it in a single entry only, they will be excluded from the tender or not being able to meet all requirements. If the supplier states „YES“ and it will be found out that the offered goods do not comply with this requirement, they can be excluded for non-compliance and for breaking the tender terms. If the tenderer fails to choose „YES“ or „NO“ where required, they can be excluded from the tenderer for not meeting the tender terms. In column „Parameter – offered value“ they state a specific parameter (in units as required) or some further specification offered performance is considered to correspond exactly to the requirement of the contracting authority specific in column „Required parameter“. The tenderer shall fill in the technical conditions according to the instructions state in them including the kind and type of goods. On filling in these kinds and types off goods the suppliers is bound to deliver them and their list will form an Annex to the Purchase Contract. This means that the supplier will be bound to deliver exactly what they have stated in their quote.

The supplier will supply new and unused Vibration table.

| Product trade name: | (Tenderer fills in the product trade name) |  |  |
| --- | --- | --- | --- |
| Required parameter | Parameter – required value | Meeting the requirements | Parameter – offered value |
| Minimal frequency range | 5 to 3300 Hz | YES / NO |  |
| Power SINE | min. 8 kN | YES / NO |  |
| Power RANDOM | min. 8 kN | YES / NO |  |
| Power SHOCK | min. 16 kN | YES / NO |  |
| Velocity SINE | min. 2,2 m/s | YES / NO |  |
| Velocity SHOCK | min. 2,2 m/s | YES / NO |  |
| Acceleration SINE | min. 1200 m/s2 | YES / NO |  |
| Acceleration RANDOM | min. 850 m/s2 rms | YES / NO |  |
| Acceleration SHOCK | min. 2000 m/s2 | YES / NO |  |
| Displacement | min. 51 mm pk-pk | YES / NO |  |
| Payload of the Vibration generator | min. 200 kg | YES / NO |  |
| Consumption electricity | max. 16,5 kVA | YES / NO |  |
| Armature diameter | min. 190 mm | YES / NO |  |
| Mass of the armature | max. 8 kg | YES / NO |  |
| Inserts on the armature | M8 threats | YES / NO |  |
| Guiding system of the armature | parallel slope guiding (rollers) | YES / NO |  |
| Thermal barrier of the armature | integrated thermal barrier of the armature | YES / NO |  |
| Top cover construction | allows a direct coupling with the climatic chamber without extra mounting any parts | YES / NO |  |
| Diameter of the Vibration generator top cover for connection of the climatic chamber | max. 565 mm | YES / NO |  |
| Dimensions of the Vibration generator (HxWxD) | max. 910 x 1100 x 600 | YES / NO |  |
| Measurement of armature zero position | induction | YES / NO |  |
| Alignment of the armature to zero position after test sample mounting / Head expander | automatic | YES / NO |  |
| Vibration generator body suspension | air springs | YES / NO |  |
| Cooling | Required, air cooled | YES / NO |  |
| Cooling blower | incl. silencer, placing outside on building roof in max. distance 10 m from the shaker | YES / NO |  |
| Cable length between amplifier and Vibration generator | min. 5m | YES / NO |  |
| Head Expander | Required | YES / NO |  |
| Head Expander mounting plane dimensions | 800x800 mm | YES / NO |  |
| Head Expander material | Aluminium | YES / NO |  |
| Head Expander frequency range | 0 to 1000 Hz | YES / NO |  |
| Head Expander mounting pattern | 100 mm, thread M8 | YES / NO |  |
| Control system part of vibration table delivery | Required | YES / NO |  |
| Minimum inputs channels | 4 | YES / NO |  |
| Inputs Charge, Voltage, IEPE | Required | YES / NO |  |
| Support TEDS connectivity | Required | YES / NO |  |
| Minimum output channels | 4 | YES / NO |  |
| Minimum sampling frequency | 100 kHz | YES / NO |  |
| A/D and D/A convertors rate | 32 bit | YES / NO |  |
| Minimal dynamic range | 100 dB | YES / NO |  |
| Communication interface | Ethernet | YES / NO |  |
| Software for tests profiles | SINE, RANDOM | YES / NO |  |
| Allows pre-defined test acc. to standards | Required | YES / NO |  |
| Number of accelerometers | 4 pcs | YES / NO |  |
| Accelerometer’s sensitivity | 10 mV/g | YES / NO |  |
| Temperature range of accelerometers | -55 °C to +120 °C | YES / NO |  |
| Frequency range of accelerometers | 1 to 11 000 Hz | YES / NO |  |
| Accelerometer’s cabling | Required | YES / NO |  |
| Length of the accelerometers cabling | min. 3m | YES / NO |  |
| Vibration generator weight | max. 1000 kg | YES / NO |  |
| EC/EU declaration of conformity | Required | YES / NO |  |
| Test report generation | Automatic | YES / NO |  |
| Operation Manual | Required | YES / NO |  |
| System Schematic | Required | YES / NO |  |
| Operator training | Onsite training required | YES / NO |  |

I (we) below signed       hereby honestly declare that the data stated above are true and genuine and that the supplier       shall deliver the goods exactly as stated in the technical and business conditions in their quote in case they are chosen in the subject public tender by the contracting authority.

In     on

_________________________________

Name and fiction of the authorized person

Stamp and signature of the authorized person